Autoclave 50 L

Model: SM530

Brand: Yamato

Origin: Japan

Technical Specifications:

Internal capacity: 50L-55L

System Automatic high-pressure steam sterilization or better.

Operating temperature: Sterilize: 105 to 135°C, Liquefy: 60 to 110°C or better.

Retain Temp.: 45 to 60°C.

Preheat temp.: 45 to 80°C.

Dry: 135 to 150°C.

Operating Ambient Temp.: 5 to 35°C.

Maximum pressure capacity: 0.255MPa or better.

Heating: Sterilize Pipe 1000W ×2 or better.

Drying Pipe: 110V/455W×2 or better.

Temp. controller: PID controlled by microprocessor or better.

Temp. setting / display: Touch panel or better.

Timer / Timer resolution Range: 0 or 1min to 99h59min / 1 min. or better.

Safety Device: Sterilize sensor error, sterilize SSR short circuit, dry sensor error, dry SSR short circuit, water level detection (liquid expansion method), independent chamber overheat protection, cover unlock error, over pressure protection, under pressure protection, warning about setting error in cooling water container, pressure switch (0.25MPa), pressure safety valve (0.255MPa) or more.

Power source: AC200~240V.

Accessories: 2 pcs. stainless steel mesh basket, 3 pcs. stainless steel mesh basket, Drain board x 1, drain bottle x 1, chemical indicator 1 set, filter x 1, droplet tray x 1.

Country of origin certificate should be submitted

Manufacturer Authorization must be submitted with the tender and strictly not accepted any other third-party authorization.

Warranty/Guarantee: 01 (One) years from the date of installation.

Country of Origin: USA/Japan/UK or Equivalent.