Arc/Spark Optical Emission Spectrometry (OES) Analyzer (Fe, Al & Cu)

Brand: SPECTRO

Model: SPECTROMAXx

Technical Specification:

Spectrometer type and calibration requirement: The quoted model should be a laboratory benchtop model and suitable for Fe, Al & Cu base application. Instrument should be providing with factory Calibration for Fe-Orientation, Low Alloy, Cast Iron, Al- Global, Pur Aluminum and Cu-Orientation, Pure Copper, Cu/Zn Alloy as per mentioned element mentioned in the specification. (Annexure-1)

Optical System: Should have dual optical system. UV and Air Optical system.

Focal Length: 300 mm focal length for UV optics and 400 mm focal length for Air optics.

Type of Detectors: CMOS multi detector.

Source Parameters: Plasma Generator–Digital sample excitation system for high precision.

Wavelength Range: Minimum wavelength coverage from 120 nm – 670 nm or better.

Spark Stand: Spark Stand self-cooled with minimum argon gas consumption. Spark stand design should be such that samples of various kind of geometries can be tested.

Diagnosis System: Instrument should be equipped with Auto diagnostics indicating the status of parts in the system by different colour indication thereby enabling the operator to take adequate preventive measure to prevent breakdowns as Much as possible.

Application report: Bidder should submit manufacturer application report for Fe, Al, Cu base application.

Standardization: Instrument should have ICAL 2.0 facility for quick standardization of the instrument. Same should be mentioned in the offer and technical confirmation by OEM.

Installation requirement / Other items required to install the instrument.

A. MBH/BAS make CRM sample -1 Nos for each base Fe, Al & Cu.

B. Computer & Printer. Computer should be supplied as per Spectrometer requirement.

C. 2 Nos. of high pure Argon Gas cylinder. (99.998% Ar)

D. 1 No. SS regulator for argon gas cylinder.

E. Sample preparation machine.

F. On-Line UPS with in-built isolation transformer with 30 minutes’ battery backup.

G. Portable Photometer Measurement of Nitrite Ion, Metal Ions, bacterial turbidity, protein, antioxidant activity, Amino acid, glucose, water and heavy metals analysis. Absorbance range: 0.02 or more. Wavelength range Color Slot: Red:575-660 nm, Green: 455-630 nm, Blue: 400-540 nm. Minimum Sample Volume: 30 μL.

Term and Condition: Bidder should submit OEM Letterhead confirmation for all the technical point need to submit along with the bid, without OEM confirmation technical point will not consider)

(Annexure-1)

Factory Calibration for Iron Base

| MAXx Advanced | Fe-01 |  | Fe-10 |  | Fe-20 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Fe-Orientation |  | low alloy steel |  | cast iron |  |
| Element | Min | Max | Min | Max | Min | Max |
| C | 0.0015 | 4.70 | 0.0008 | 1.50 | 1.40 | 4.70 |
| Si | 0.0005 | 20.00 | 0.0005 | 20.00 | 0.0010 | 4.80 |
| Mn | 0.0005 | 29.00 | 0.0005 | 2.30 | 0.0010 | 4.10 |
| P | 0.0005 | 2.40 | 0.0005 | 0.150 | 0.0005 | 2.40 |
| S | 0.0005 | 0.43 | 0.0005 | 0.120 | 0.0005 | 0.240 |
| Cr | 0.0010 | 33.0 | 0.0003 | 8.40 | 0.0003 | 2.50 |
| Mo | 0.0005 | 9.50 | 0.0004 | 8.50 | 0.0004 | 2.20 |
| Ni | 0.0010 | 46.0 | 0.0005 | 5.50 | 0.0005 | 5.40 |
| Al | 0.0005 | 8.70 | 0.0005 | 2.00 | 0.0005 | 1.40 |
| Co | 0.0010 | 18.00 | 0.0006 | 2.10 | 0.0006 | 0.210 |
| Cu | 0.0005 | 8.10 | 0.0002 | 1.20 | 0.0003 | 2.50 |
| Nb | 0.0010 | 3.00 | 0.0005 | 0.56 | 0.0005 | 0.280 |
| Ti | 0.0005 | 3.00 | 0.0002 | 0.82 | 0.0002 | 0.45 |
| V | 0.0005 | 10.00 | 0.0005 | 1.10 | 0.0005 | 0.62 |
| W | 0.0050 | 21.00 | 0.0050 | 3.10 | 0.0050 | 0.130 |
| Pb | 0.0010 | 0.380 | 0.0008 | 0.0250 | 0.0008 | 0.060 |
| Sn | 0.0005 | 0.240 | 0.0002 | 0.140 | 0.0002 | 0.240 |
| Mg | 0.0010 | 0.260 |  |  | 0.0010 | 0.130 |
| As | 0.0010 | 0.200 | 0.0010 | 0.140 | 0.0010 | 0.200 |
| Zr | 0.0010 | 0.350 | 0.0010 | 0.350 | 0.0010 | 0.070 |
| Bi | 0.0010 | 0.067 | 0.0010 | 0.067 | 0.0010 | 0.0210 |
| Ca | 0.0001 | 0.0130 | 0.0001 | 0.0130 | 0.0001 | 0.0060 |
| Ce | 0.0010 | 0.170 | 0.0010 | 0.170 | 0.0010 | 0.100 |
| Sb | 0.0010 | 0.240 | 0.0010 | 0.120 | 0.0010 | 0.240 |
| Se | 0.0020 | 0.310 | 0.0010 | 0.130 | 0.0015 | 0.060 |
| Te | 0.0010 | 0.080 | 0.0010 | 0.0210 | 0.0010 | 0.080 |
| Ta | 0.0100 | 0.76 | 0.0070 | 0.210 |  |  |
| B | 0.0002 | 2.00 | 0.0001 | 0.0230 | 0.0001 | 0.110 |
| Zn | 0.0010 | 0.045 | 0.0005 | 0.0260 | 0.0007 | 0.045 |
| La | 0.0005 | 0.045 | 0.0003 | 0.045 | 0.0003 | 0.0260 |
| Ag | 0.0001 | 0.0050 | 0.00005 | 0.0050 |  |  |
| Nd | 0.0010 | 0.0200 |  |  | 0.0010 | 0.0200 |

Factory Calibration for Aluminum Base

| MAXx Advanced | Al-01 |  | Al-10 |  |
| --- | --- | --- | --- | --- |
|  | Al - Globalcalibration |  | Al - pure/low alloys |  |
| Element | Min | Max | Min | Max |
| Si | 0.0005 | 36.0 | 0.0005 | 1.90 |
| Fe | 0.0005 | 12.00 | 0.0005 | 6.10 |
| Cu | 0.0001 | 53.0 | 0.0001 | 1.10 |
| Mn | 0.0002 | 30.0 | 0.0002 | 1.20 |
| Mg | 0.00005 | 11.00 | 0.00005 | 1.20 |
| Cr | 0.0002 | 0.46 | 0.0002 | 0.310 |
| Ni | 0.0003 | 4.00 | 0.0003 | 1.10 |
| Zn | 0.0008 | 20.00 | 0.0008 | 1.10 |
| Ti | 0.0003 | 5.40 | 0.0003 | 0.220 |
| Ag | 0.0001 | 1.00 | 0.0001 | 0.100 |
| As | 0.0010 | 0.0110 | 0.0010 | 0.0080 |
| B | 0.0001 | 1.50 | 0.0001 | 0.0220 |
| Ba | 0.0001 | 0.0060 | 0.0001 | 0.0060 |
| Be | 0.00005 | 0.0210 | 0.00005 | 0.0210 |
| Bi | 0.0005 | 0.71 | 0.0005 | 0.160 |
| Ca | 0.0001 | 0.040 | 0.0001 | 0.041 |
| Cd | 0.0001 | 0.330 | 0.0001 | 0.210 |
| Ce | 0.0010 | 0.046 | 0.0010 | 0.046 |
| Co | 0.0004 | 1.60 | 0.0004 | 0.100 |
| Ga | 0.0002 | 0.110 | 0.0001 | 0.110 |
| Hg | 0.0003 | 0.0100 | 0.0003 | 0.0100 |
| In | 0.0003 | 0.100 | 0.0003 | 0.100 |
| La | 0.0003 | 11.00 | 0.0003 | 0.0310 |
| Li | 0.0001 | 8.20 | 0.0001 | 0.0250 |
| Mo | 0.0005 | 0.120 | 0.0003 | 0.0120 |
| Na | 0.0001 | 0.0210 | 0.0001 | 0.0210 |
| P | 0.0008 | 0.0170 | 0.0008 | 0.0100 |
| Pb | 0.0005 | 1.50 | 0.0005 | 1.40 |
| Sb | 0.0015 | 0.55 | 0.0015 | 0.310 |
| Sn | 0.0003 | 21.00 | 0.0002 | 1.10 |
| Sr | 0.00005 | 11.00 | 0.00005 | 0.0100 |
| Tl |  |  | 0.0005 | 0.0150 |
| V | 0.0005 | 0.110 | 0.0005 | 0.110 |
| Zr | 0.0002 | 1.30 | 0.0002 | 0.110 |
| Hf | 0.0030 | 10.00 |  |  |
| Sc | 0.0002 | 0.40 |  |  |
| Y | 0.0003 | 0.50 |  |  |

Factory Calibration for Copper Base

| MAXx Advanced | Cu-01 |  | Cu-10 |  | Cu-20 |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Cu-Orientation |  | pure copper |  | Cu/Zn - alloy |  |
| Element | Min | Max | Min | Max | Min | Max |
| Zn | 0.0015 | 50.0 | 0.0005 | 0.300 | 0.80 | 50.0 |
| Pb | 0.0010 | 22.00 | 0.0003 | 1.50 | 0.0010 | 5.10 |
| Sn | 0.0005 | 14.80 | 0.0002 | 0.330 | 0.0002 | 9.40 |
| P | 0.0005 | 1.10 | 0.0002 | 0.085 | 0.0005 | 0.250 |
| Mn | 0.0002 | 24.00 | 0.0001 | 0.120 | 0.0003 | 18.50 |
| Fe | 0.0002 | 6.60 | 0.0002 | 0.220 | 0.0005 | 4.60 |
| Ni | 0.0005 | 34.0 | 0.0002 | 0.45 | 0.0005 | 4.50 |
| Si | 0.0005 | 7.15 | 0.0005 | 0.090 | 0.0005 | 7.10 |
| Mg | 0.0001 | 0.180 | 0.0001 | 0.0200 | 0.0001 | 0.0210 |
| Cr | 0.0002 | 2.40 | 0.0002 | 0.95 | 0.0002 | 0.090 |
| Te | 0.0003 | 0.120 | 0.0003 | 0.120 | 0.0005 | 0.0120 |
| As | 0.0004 | 0.78 | 0.0002 | 0.330 | 0.0002 | 0.220 |
| Sb | 0.0035 | 1.70 | 0.0008 | 0.350 | 0.0030 | 0.80 |
| Cd | 0.0001 | 0.110 | 0.0001 | 0.110 | 0.0002 | 0.0350 |
| Bi | 0.0005 | 5.80 | 0.0005 | 0.070 | 0.0005 | 5.80 |
| Ag | 0.0003 | 2.10 | 0.0002 | 0.53 | 0.0010 | 0.041 |
| Co | 0.0015 | 2.40 | 0.0003 | 0.200 | 0.0010 | 0.350 |
| Al | 0.0005 | 12.50 | 0.0003 | 0.040 | 0.0005 | 10.70 |
| S | 0.0002 | 0.220 | 0.0002 | 0.080 | 0.0002 | 0.050 |
| Be | 0.0001 | 3.00 | 0.00005 | 0.0210 | 0.0001 | 0.0100 |
| Zr | 0.0002 | 0.380 | 0.0002 | 0.230 |  |  |
| Au | 0.0005 | 0.100 | 0.0005 | 0.100 |  |  |
| B | 0.0001 | 0.0350 | 0.0001 | 0.0350 | 0.0002 | 0.0060 |
| C | 0.0005 | 0.061 |  |  |  |  |
| Ti | 0.0002 | 0.78 | 0.0002 | 0.0080 |  |  |
| Se | 0.0008 | 1.70 | 0.0002 | 0.120 | 0.0008 | 1.70 |
| Nb | 0.0010 | 1.30 |  |  |  |  |
| Li |  |  | 0.00005 | 0.0130 |  |  |
| Pd |  |  | 0.0010 | 0.0250 |  |  |
| Pt | 0.0020 | 0.0220 | 0.0020 | 0.0220 |  |  |
| Ir |  |  | 0.0010 | 0.0250 |  |  |